# CardioLens Architecture Diagram Specification

## Device Overview: Quad-Modal Cardiac Imaging System

## 1. Physical Form Factor

TOP VIEW (Patient-facing surface)

                         ←── 80mm ──→
                    ╭─────────────────────╮
                   ╱                       ╲
                  │    ┌───────────────┐    │
                  │    │   NMR COIL    │    │     ↑
                  │    │   (Primary)   │    │     │
            ↑     │    │   Ø50mm      │    │     │
            │     │    └───────────────┘    │    100mm
           60mm   │  ○ ○ ○ ○ ○ ○ ○ ○ ○ ○   │     │
            │     │  ECG Dry Electrodes     │     │
            ↓     │  (Gold-polymer, 12x)    │     │
                  │                         │     ↓
                   ╲    [PPG] [MIC] [PPG]  ╱
                    ╰─────────────────────╯
                         Bio-adhesive rim

                    Curvature: R = 150mm (conformal)

### Dimensions

| Parameter | Value | Rationale |
| --- | --- | --- |
| Length | 100mm | Spans intercostal spaces |
| Width | 80mm | Fits between ribs |
| Thickness | 15mm (center) | Houses Halbach array |
| Curvature radius | 150mm | Matches average adult thorax |
| Weight target | <200g | Comparable to smartphone |

## 2. Layer Stack (Cross-Section)

SIDE VIEW (Cross-section through center)

    Skin surface ════════════════════════════════════════
                 ↓ Contact layer (bio-adhesive polymer)
    ┌────────────────────────────────────────────────────┐
    │ ▓▓▓▓▓▓▓▓▓ Silicone cushion (2mm) ▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓ │
    ├────────────────────────────────────────────────────┤
    │ ○ ○ ○ ○ ○ Dry ECG electrodes (Au/polymer) ○ ○ ○ ○ │ ← Perimeter
    ├────────────────────────────────────────────────────┤
    │ ┌──────────────────────────────────────────────┐   │
    │ │  RF Coil Layer (flexible PCB, polyimide)     │   │ ← 0.5mm
    │ │  Transmit/Receive switchable                 │   │
    │ └──────────────────────────────────────────────┘   │
    ├────────────────────────────────────────────────────┤
    │ ┌─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┐   │
    │ │N│S│N│S│N│S│N│S│N│S│N│S│N│S│N│S│N│S│N│S│N│S│N│   │ ← Halbach Array
    │ │↓│→│↑│←│↓│→│↑│←│↓│→│↑│←│↓│→│↑│←│↓│→│↑│←│↓│→│↑│   │   (NdFeB, 8mm)
    │ └─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┘   │
    ├────────────────────────────────────────────────────┤
    │ ┌──────────────────────────────────────────────┐   │
    │ │  Electronics PCB (rigid-flex, 4-layer)       │   │ ← 2mm
    │ │  • ADC (24-bit, 4-channel)                   │   │
    │ │  • MCU (Cortex-M4F)                          │   │
    │ │  • RF frontend (LNA, mixer, diplexer)        │   │
    │ │  • Bluetooth LE + WiFi                       │   │
    │ └──────────────────────────────────────────────┘   │
    ├────────────────────────────────────────────────────┤
    │ ▓▓▓▓▓▓▓▓▓ Battery (Li-Po, 500mAh) ▓▓▓▓▓▓▓▓▓▓▓▓▓▓ │ ← 2mm
    ├────────────────────────────────────────────────────┤
    │ ═════════ Enclosure (medical-grade ABS) ═════════ │ ← 0.5mm
    └────────────────────────────────────────────────────┘

## 3. Sensing Modality Integration

### 3.1 Spatial Distribution

COMPONENT PLACEMENT (Top-down, exploded)

    Layer 1: Contact Surface
    ╭─────────────────────────────────────╮
    │  ○   ○   ○   ○   ○   ○   ○   ○   ○ │  ← ECG electrodes (perimeter)
    │ ○                               ○   │
    │     ┌─────────────────────────┐     │
    │ ○   │                         │  ○  │
    │     │    RF COIL APERTURE     │     │
    │ ○   │       (clear)           │  ○  │
    │     │                         │     │
    │ ○   └─────────────────────────┘  ○  │
    │                                     │
    │ ○  [PPG-L]    [MIC]    [PPG-R]   ○  │  ← Optical + Acoustic
    │  ○   ○   ○   ○   ○   ○   ○   ○   ○ │
    ╰─────────────────────────────────────╯

    PPG-L/R: Dual PPG for differential measurement
    MIC: MEMS microphone (acoustic window)

### 3.2 Signal Routing

SIGNAL FLOW DIAGRAM

    ┌─────────────┐
    │   SENSORS   │
    └──────┬──────┘
           │
    ┌──────┴──────────────────────────────────────────────┐
    │                                                      │
    ▼                   ▼                 ▼               ▼
┌───────┐         ┌─────────┐       ┌─────────┐     ┌─────────┐
│  ECG  │         │   NMR   │       │   PPG   │     │ ACOUSTIC│
│12 elec│         │RF coil  │       │LED+PD×2 │     │MEMS mic │
└───┬───┘         └────┬────┘       └────┬────┘     └────┬────┘
    │                  │                 │               │
    ▼                  ▼                 ▼               ▼
┌───────┐         ┌─────────┐       ┌─────────┐     ┌─────────┐
│ Inst  │         │  LNA    │       │Trans-   │     │ Pre-amp │
│ Amp   │         │ 0.5dB NF│       │impedance│     │ + AGC   │
│×1000  │         └────┬────┘       │ Amp     │     └────┬────┘
└───┬───┘              │            └────┬────┘          │
    │                  ▼                 │               │
    │            ┌─────────┐             │               │
    │            │ Diplexer│             │               │
    │            │PIN diode│             │               │
    │            │ switch  │             │               │
    │            └────┬────┘             │               │
    │                 │                  │               │
    ▼                 ▼                  ▼               ▼
┌─────────────────────────────────────────────────────────┐
│                    24-bit ADC (4 channels)              │
│              Simultaneous sampling @ 10 kHz             │
│         (ECG: 1kHz, NMR: variable, PPG: 100Hz,         │
│                    Acoustic: 8kHz)                      │
└────────────────────────────┬────────────────────────────┘
                             │
                             ▼
                    ┌─────────────────┐
                    │   MCU (M4F)     │
                    │ Signal fusion   │
                    │ ML inference    │
                    │ Data buffering  │
                    └────────┬────────┘
                             │
              ┌──────────────┴──────────────┐
              ▼                              ▼
      ┌─────────────┐                ┌─────────────┐
      │ Bluetooth LE│                │   Storage   │
      │ (real-time) │                │  (32GB μSD) │
      └──────┬──────┘                └─────────────┘
             │
             ▼
      ┌─────────────┐
      │ Companion   │
      │ App (iOS/   │
      │ Android)    │
      └─────────────┘

## 4. Halbach Array Configuration

### 4.1 Magnetic Field Pattern

HALBACH ARRAY CROSS-SECTION (Magnetic Field Lines)

    Weak field side (back of device)
    ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─ ─

    ┌─────┬─────┬─────┬─────┬─────┬─────┬─────┬─────┐
    │  ↓  │  →  │  ↑  │  ←  │  ↓  │  →  │  ↑  │  ←  │  Magnet
    │  N  │ N→S │  S  │ S←N │  N  │ N→S │  S  │ S←N │  Orientation
    └─────┴─────┴─────┴─────┴─────┴─────┴─────┴─────┘

           ╲   │   ╱       ╲   │   ╱
            ╲  │  ╱         ╲  │  ╱
             ╲ │ ╱           ╲ │ ╱         Field lines
              ╲│╱             ╲│╱          (concentrated
               │               │           on patient side)
               │               │
    ═══════════════════════════════════════════════════
    Strong field side (toward patient/heart)

    B₀ ≈ 0.05-0.1 T at coil surface
    Penetration: 30-40mm into tissue

### 4.2 Field Strength Profile

AXIAL FIELD DECAY (z = depth into tissue)

    B(z)
    (mT)
    │
 100├●
    │ ╲
  80├  ●
    │   ╲
  60├    ●
    │     ╲
  40├      ●──────── Target imaging depth
    │        ╲        (left ventricle)
  20├         ●
    │          ╲
   0├───────────●────●────●─────→ z (mm)
    0    10    20   30   40   50

    B(z) = μ₀IR² / 2(R² + z²)^(3/2)

    For R = 25mm, I = 2A:
    • z = 0mm:  B = 100 mT
    • z = 20mm: B = 45 mT
    • z = 40mm: B = 15 mT

## 5. ECG Electrode Configuration

12-LEAD EQUIVALENT PLACEMENT

    Standard chest positions mapped to device footprint:

              V1   V2
               ○   ○

         ○               ○

     ○       ┌─────┐       ○
             │ NMR │
         ○   │COIL │   ○
             └─────┘
         ○               ○

               ○   ○
              V5   V6

    Differential pairs for noise rejection:
    • V1-V2: Septal
    • V3-V4: Anterior
    • V5-V6: Lateral

    Electrode specs:
    • Material: Au-plated polymer (dry contact)
    • Diameter: 8mm each
    • Contact impedance: <50 kΩ @ 10 Hz

## 6. Optical (PPG) Subsystem

PPG MODULE DETAIL

    ┌─────────────────────────────────────┐
    │                                     │
    │   ┌─────┐       ┌─────┐            │
    │   │ LED │       │ LED │            │
    │   │ 530 │       │ 940 │            │
    │   │ nm  │       │ nm  │            │
    │   │green│       │ IR  │            │
    │   └──┬──┘       └──┬──┘            │
    │      │             │               │
    │      ▼             ▼               │
    │   ━━━━━━━━━━━━━━━━━━━━━  Skin     │
    │   ┬─┬─┬─┬─┬─┬─┬─┬─┬─┬─┬  Tissue   │
    │   │●│●│●│●│●│●│●│●│●│●│  Blood    │
    │   ┴─┴─┴─┴─┴─┴─┴─┴─┴─┴─┴           │
    │         │ Reflected                │
    │         ▼ light                    │
    │      ┌─────┐                       │
    │      │ PD  │  Photodiode           │
    │      │     │  (Si, filtered)       │
    │      └─────┘                       │
    │                                     │
    └─────────────────────────────────────┘

    Measurements:
    • SpO₂ (blood oxygen saturation)
    • Heart rate (pulse detection)
    • Pulse waveform morphology
    • Perfusion index

## 7. Acoustic Subsystem

MEMS MICROPHONE INTEGRATION

    Cross-section:

    Device enclosure
    ═══════════════════════════════════
              ┌─────────┐
              │░░░░░░░░░│  Acoustic mesh
              │░░░░░░░░░│  (dust/moisture
              └────┬────┘   protection)
                   │
              ┌────┴────┐
              │  ┌───┐  │
              │  │MIC│  │  MEMS microphone
              │  └─┬─┘  │  • Sensitivity: -38 dBFS
              │    │    │  • SNR: 65 dB
              │  ┌─┴─┐  │  • Freq range: 20-20kHz
              │  │ADC│  │  (cardiac: 20-600 Hz)
              │  └───┘  │
              └─────────┘

    Target sounds:
    • S1 (lub): Mitral/tricuspid closure, 25-45 Hz
    • S2 (dub): Aortic/pulmonic closure, 50-80 Hz
    • Murmurs: Turbulent flow, 100-600 Hz
    • S3/S4 (gallops): Ventricular filling, 25-50 Hz

## 8. System Block Diagram

┌────────────────────────────────────────────────────────────────────────┐
│                           CARDIOLENS SYSTEM                            │
├────────────────────────────────────────────────────────────────────────┤
│                                                                        │
│  ┌──────────────────────────────────────────────────────────────────┐  │
│  │                        SENSING LAYER                              │  │
│  │                                                                    │  │
│  │  ┌────────┐  ┌────────┐  ┌────────┐  ┌────────┐                   │  │
│  │  │  ECG   │  │  NMR   │  │  PPG   │  │ACOUSTIC│                   │  │
│  │  │        │  │        │  │        │  │        │                   │  │
│  │  │ 12 dry │  │ RF coil│  │ Dual   │  │ MEMS   │                   │  │
│  │  │ elec-  │  │ Halbach│  │ wave-  │  │ micro- │                   │  │
│  │  │ trodes │  │ array  │  │ length │  │ phone  │                   │  │
│  │  └───┬────┘  └───┬────┘  └───┬────┘  └───┬────┘                   │  │
│  │      │           │           │           │                        │  │
│  └──────┼───────────┼───────────┼───────────┼────────────────────────┘  │
│         │           │           │           │                           │
│  ┌──────┼───────────┼───────────┼───────────┼────────────────────────┐  │
│  │      ▼           ▼           ▼           ▼    ANALOG FRONTEND     │  │
│  │  ┌────────┐  ┌────────┐  ┌────────┐  ┌────────┐                   │  │
│  │  │ Instr  │  │  LNA   │  │ Trans- │  │ Pre-   │                   │  │
│  │  │ Amp    │  │ +RF FE │  │ imped  │  │ amp    │                   │  │
│  │  │ CMRR   │  │ Diplex │  │ Amp    │  │ +AGC   │                   │  │
│  │  │ >100dB │  │ Switch │  │        │  │        │                   │  │
│  │  └───┬────┘  └───┬────┘  └───┬────┘  └───┬────┘                   │  │
│  │      │           │           │           │                        │  │
│  │      ▼           ▼           ▼           ▼                        │  │
│  │  ┌───────────────────────────────────────────────────────────┐    │  │
│  │  │              24-bit Sigma-Delta ADC (4 channels)          │    │  │
│  │  │                  Simultaneous sampling                     │    │  │
│  │  └─────────────────────────┬─────────────────────────────────┘    │  │
│  │                            │                                      │  │
│  └────────────────────────────┼──────────────────────────────────────┘  │
│                               │                                         │
│  ┌────────────────────────────┼──────────────────────────────────────┐  │
│  │                            ▼         DIGITAL PROCESSING           │  │
│  │  ┌─────────────────────────────────────────────────────────────┐  │  │
│  │  │                    MCU (ARM Cortex-M4F)                      │  │  │
│  │  │  ┌─────────┐ ┌─────────┐ ┌─────────┐ ┌─────────┐            │  │  │
│  │  │  │  ECG    │ │  NMR    │ │  PPG    │ │ Acoustic│            │  │  │
│  │  │  │ Process │ │ Process │ │ Process │ │ Process │            │  │  │
│  │  │  │ •Filter │ │ •FFT    │ │ •SpO2   │ │ •Filter │            │  │  │
│  │  │  │ •QRS    │ │ •T1/T2  │ │ •HR     │ │ •S1/S2  │            │  │  │
│  │  │  │ •ST seg │ │ •Recon  │ │ •Pulse  │ │ •Murmur │            │  │  │
│  │  │  └────┬────┘ └────┬────┘ └────┬────┘ └────┬────┘            │  │  │
│  │  │       │           │           │           │                  │  │  │
│  │  │       └───────────┴─────┬─────┴───────────┘                  │  │  │
│  │  │                         │                                    │  │  │
│  │  │                         ▼                                    │  │  │
│  │  │               ┌─────────────────┐                            │  │  │
│  │  │               │  SENSOR FUSION  │                            │  │  │
│  │  │               │   ML Engine     │                            │  │  │
│  │  │               │  (TinyML/CMSIS) │                            │  │  │
│  │  │               └────────┬────────┘                            │  │  │
│  │  │                        │                                     │  │  │
│  │  └────────────────────────┼─────────────────────────────────────┘  │  │
│  │                           │                                        │  │
│  │            ┌──────────────┼──────────────┐                         │  │
│  │            ▼              ▼              ▼                         │  │
│  │     ┌───────────┐  ┌───────────┐  ┌───────────┐                   │  │
│  │     │ BLE 5.0   │  │ Storage   │  │ Display   │                   │  │
│  │     │ Real-time │  │ microSD   │  │ LEDs      │                   │  │
│  │     │ streaming │  │ logging   │  │ Status    │                   │  │
│  │     └───────────┘  └───────────┘  └───────────┘                   │  │
│  │                                                                    │  │
│  └────────────────────────────────────────────────────────────────────┘  │
│                                                                          │
│  ┌────────────────────────────────────────────────────────────────────┐  │
│  │                        POWER MANAGEMENT                            │  │
│  │  ┌────────────┐    ┌────────────┐    ┌────────────┐               │  │
│  │  │ Li-Po      │───▶│ PMIC       │───▶│ Regulators │               │  │
│  │  │ 500mAh     │    │ Charging   │    │ 3.3V, 5V   │               │  │
│  │  │ USB-C chg  │    │ Protection │    │ RF supply  │               │  │
│  │  └────────────┘    └────────────┘    └────────────┘               │  │
│  └────────────────────────────────────────────────────────────────────┘  │
│                                                                          │
└────────────────────────────────────────────────────────────────────────┘

## 9. Signal Timing Diagram

MULTI-MODAL ACQUISITION TIMING

    Time →  0ms    100ms   200ms   300ms   400ms   500ms   600ms
            │       │       │       │       │       │       │
    ECG     ▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓
            Continuous @ 1 kHz sampling

    NMR     ██░░░░░░██░░░░░░██░░░░░░██░░░░░░██░░░░░░██░░░░░░
            50ms bursts @ 5 Hz repetition
            (RF off during other acquisitions to avoid interference)

    PPG     █░░░░░░░█░░░░░░░█░░░░░░░█░░░░░░░█░░░░░░░█░░░░░░░
            Triggered @ 100 Hz, 1ms LED pulse
            (Interleaved with NMR quiet periods)

    Acoustic▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓▓
            Continuous @ 8 kHz sampling
            (Filtered to remove RF artifacts during NMR bursts)

    Legend: ▓ = Continuous  █ = Burst  ░ = Idle/blanking

    Key constraints:
    • NMR and PPG LED do not overlap (RF interference)
    • ECG and Acoustic run continuously (different frequency domains)
    • 200ms window captures full cardiac cycle (~300ms at 60 bpm)

## 10. Data Fusion Architecture

MULTI-MODAL FUSION PIPELINE

    Raw Signals                   Feature Extraction              Fusion
    ───────────                   ──────────────────              ──────

    ECG ────────────┐    ┌───────────────────────────────┐
    (1 kHz)         │    │ • QRS complex detection       │
                    ├───▶│ • ST segment analysis         │
                    │    │ • Heart rate variability      │
                    │    │ • Arrhythmia classification   │───┐
                    │    └───────────────────────────────┘   │
                    │                                        │
    NMR ────────────┤    ┌───────────────────────────────┐   │
    (5 Hz bursts)   │    │ • T1/T2 relaxation times      │   │
                    ├───▶│ • Tissue characterization     │   │    ┌─────────┐
                    │    │ • Edema detection             │───┼───▶│  MULTI- │
                    │    │ • Fibrosis markers            │   │    │  MODAL  │
                    │    └───────────────────────────────┘   │    │ CARDIAC │
                    │                                        │    │  INDEX  │
    PPG ────────────┤    ┌───────────────────────────────┐   │    │  (MCI)  │
    (100 Hz)        │    │ • SpO₂ calculation            │   │    └────┬────┘
                    ├───▶│ • Pulse transit time          │───┤         │
                    │    │ • Perfusion index             │   │         │
                    │    │ • Waveform morphology         │   │         ▼
                    │    └───────────────────────────────┘   │    ┌─────────┐
                    │                                        │    │ CLINICAL│
    Acoustic ───────┤    ┌───────────────────────────────┐   │    │ DECISION│
    (8 kHz)         │    │ • S1/S2 detection             │   │    │ SUPPORT │
                    └───▶│ • Murmur classification       │───┘    │         │
                         │ • Valve timing                │        │ Risk:   │
                         │ • Extra heart sounds (S3/S4)  │        │ LOW /   │
                         └───────────────────────────────┘        │ MEDIUM /│
                                                                  │ HIGH    │
                                                                  └─────────┘

## 11. Safety Architecture

SAFETY INTERLOCK SYSTEM

    ┌─────────────────────────────────────────────────────────────────┐
    │                       SAFETY MONITORS                           │
    │                                                                 │
    │  ┌──────────┐  ┌──────────┐  ┌──────────┐  ┌──────────┐       │
    │  │Temperature│  │ Current  │  │  Time    │  │ Skin     │       │
    │  │ Sensor   │  │ Monitor  │  │ Watchdog │  │ Contact  │       │
    │  │ <41°C    │  │ <100mA   │  │ 30min max│  │ Detect   │       │
    │  └────┬─────┘  └────┬─────┘  └────┬─────┘  └────┬─────┘       │
    │       │             │             │             │              │
    │       └──────┬──────┴──────┬──────┴──────┬──────┘              │
    │              │             │             │                     │
    │              ▼             ▼             ▼                     │
    │         ┌────────────────────────────────────┐                 │
    │         │        SAFETY LOGIC (HW)           │                 │
    │         │     Independent of main MCU        │                 │
    │         │     Hardwired shutdown paths       │                 │
    │         └─────────────────┬──────────────────┘                 │
    │                           │                                    │
    │              ┌────────────┼────────────┐                      │
    │              ▼            ▼            ▼                      │
    │        ┌─────────┐  ┌─────────┐  ┌─────────┐                  │
    │        │ RF OFF  │  │POWER OFF│  │ ALERT   │                  │
    │        │ Switch  │  │ Relay   │  │ USER    │                  │
    │        └─────────┘  └─────────┘  └─────────┘                  │
    │                                                                │
    └─────────────────────────────────────────────────────────────────┘

    SAR Limit: < 2 W/kg (FDA/IEC 60601 compliant)
    Temperature rise: < 1°C surface
    Battery: UL2054 certified cells with PTC protection

## 12. Companion App Interface

MOBILE APP SCREENS (Wireframe)

    ┌─────────────────┐   ┌─────────────────┐   ┌─────────────────┐
    │    HOME         │   │   LIVE VIEW     │   │   RESULTS       │
    │                 │   │                 │   │                 │
    │  ┌───────────┐  │   │  ECG Trace      │   │  Risk Score     │
    │  │ CardioLens│  │   │  ∿∿∿∿∿∿∿∿∿∿    │   │  ┌───────────┐  │
    │  │   One     │  │   │                 │   │  │    72     │  │
    │  │    🔋     │  │   │  SpO₂: 98%      │   │  │   /100    │  │
    │  │   85%     │  │   │  HR: 72 bpm     │   │  │   LOW     │  │
    │  └───────────┘  │   │                 │   │  └───────────┘  │
    │                 │   │  ┌───────────┐  │   │                 │
    │  [Start Scan]   │   │  │ ♡ Heart   │  │   │  ECG: Normal    │
    │                 │   │  │  Sounds   │  │   │  NMR: Normal    │
    │  Last scan:     │   │  │ S1──S2──  │  │   │  SpO₂: 98%      │
    │  2 hours ago    │   │  └───────────┘  │   │  Sounds: Clear  │
    │                 │   │                 │   │                 │
    │  History  >     │   │  [◉ Recording] │   │  [Share Report] │
    │  Settings >     │   │                 │   │  [View Details] │
    │                 │   │  00:45 elapsed  │   │                 │
    └─────────────────┘   └─────────────────┘   └─────────────────┘

## 13. Manufacturing Notes

### PCB Stack-up

Layer 1: Signal (ECG traces, sensitive analog)
    ────────────────────────────────────────────────
    Prepreg (0.2mm)
    ────────────────────────────────────────────────
    Layer 2: Ground plane (solid copper)
    ────────────────────────────────────────────────
    Core (0.8mm)
    ────────────────────────────────────────────────
    Layer 3: Power planes (+3.3V, +5V, RF)
    ────────────────────────────────────────────────
    Prepreg (0.2mm)
    ────────────────────────────────────────────────
    Layer 4: Digital (MCU, BLE, connectors)

    Total thickness: ~1.6mm
    Material: FR-4 (rigid) + Polyimide (flex sections)

### Bill of Materials (Key Components)

| Component | Part Number (Example) | Qty |
| --- | --- | --- |
| MCU | STM32L4R5 | 1 |
| ADC | ADS1298 (24-bit, 8ch) | 1 |
| LNA | MAX2640 | 1 |
| PPG AFE | MAX86141 | 1 |
| MEMS Mic | ICS-43434 | 1 |
| BLE Module | nRF52840 | 1 |
| Battery | LP503035 (500mAh) | 1 |
| Halbach magnets | N52 NdFeB (custom) | 24 |

## 14. Regulatory Pathway

FDA CLASSIFICATION AND PATHWAY

    ┌─────────────────────────────────────────────────────────────────┐
    │                                                                 │
    │   Device: CardioLens (Multi-modal Cardiac Monitor)              │
    │                                                                 │
    │   Classification: Class II (510(k) pathway)                     │
    │                                                                 │
    │   Predicate devices:                                            │
    │   • ECG monitors (e.g., AliveCor KardiaMobile)                 │
    │   • Pulse oximeters (e.g., Masimo MightySat)                   │
    │   • Digital stethoscopes (e.g., Eko DUO)                       │
    │   • Portable MRI (Hyperfine Swoop - for technology principle)  │
    │                                                                 │
    │   Key standards:                                                │
    │   • IEC 60601-1 (Medical electrical equipment)                 │
    │   • IEC 60601-2-25 (ECG specific)                              │
    │   • IEC 60601-2-33 (MR equipment specific)                     │
    │   • ISO 80601-2-61 (Pulse oximeter)                            │
    │                                                                 │
    │   Testing requirements:                                         │
    │   • EMC testing (IEC 60601-1-2)                                │
    │   • Biocompatibility (ISO 10993)                               │
    │   • SAR testing (IEC 62209)                                    │
    │   • Software validation (IEC 62304)                            │
    │                                                                 │
    └─────────────────────────────────────────────────────────────────┘

Diagram specification prepared for CardioLens (Cardiac-NSIM) project. These specifications are intended for engineering visualization, CAD modeling, and regulatory documentation. Created with assistance from Claude Code (Anthropic).